## statuses
dev_id | dev_desc | bit | bit_value | status_flag | auto_status_flag | qc_flag | bit_desc | Feedback | Replies | Workspace
1 | Palas Fidas 100 | 16 | 1 | 1 | 2 | 9 | Sensor flow error | The sensor flow error is triggered by the flow and the flow velocity. The error flag is set if the flow is 5% off or if the velocity is 15% off the calibrated value. | Data is likely invalid if this status flag is raied
1 | Palas Fidas 100 | 17 | 1 | 1 | 2 | 9 | Particle coincidence | up to 20% coincidence ignored, > 20% coincidence would result from only very high concentrations and likely to be invalid (seen once in field with a spider silk flicking around the light chamber) |  | status_flag | status_desc
1 | Palas Fidas 100 | 18 | 1 | 1 | 1 | 2 | Suction pump error | flow, velocity and pump suction are all related, need to check the values of all three to determine | Data is valid if flow is OK | 0 | valid
1 | Palas Fidas 100 | 19 | 1 | 1 | 1 | 2 | Weather station error | P, T, rh - used to calc volumetric flow from mass flow, without weather we revert to mass flow based on 20 C/1 atm, result could be a small error in calc of mass concentration (3-7%) | Instrument reverts to default T&P as explaiend when Weather station sensors are not available or faulty. | 1 | may be valid
1 | Palas Fidas 100 | 20 | 1 | 1 | 0 | 0 | IADS error | indication of fogging/failure, concern is measuring water particles in addtion to pollution, but would see particle coincidence fail as well. We are not heating incoming air, IADS is not connected, should not be an error, if there is , ignore it |  | 2 | invalid
1 | Palas Fidas 100 | 21 | 1 | 1 | 1 | 2 | Calibration | +/- 3.5 channel (~6% mass error, worst case) variation allowed, generally drifts down, if this flag occured would need to recalibrate, data within instrument uncertainty |  | 4 | warm up
1 | Palas Fidas 100 | 22 | 1 | 1 | 2 | 9 | LED temperature | provides stability of temperature measurement, compensates for ambient temperature, high temp indicates power failure, low temp or off indicates everything failing - there will be no data |  | 7 | high humidity
1 | Palas Fidas 100 | 23 | 1 | 2 | 2 | 9 | Operating mode not auto | 0 = warmed up, stable, running as it should, 1s and 2s during a repower, warming up, 3 during calibration (i.e. recalibrating)
1 | Palas Fidas 100 | 29 | 1 | 2 | 2 | 9 | Mode of operation 1: idle or 3: calibration (see bit 30) |  |  | 2 | 1
1 | Palas Fidas 100 | 30 | 1 | 2 | 2 | 9 | Mode of operation 3: calibration |  |  | 30 | 29 | mode of operation
2 | Naneos Partector | 0 | 1 | 2 | 2 | 9 | Pulse low error: non-zero Charging current when high voltage is off (contamination in the charger) | yes, critical error |  | 0 | 0 | 0 | auto
2 | Naneos Partector | 1 | 1 | 2 | 2 | 9 | Pulse high error: Setpoint diffusion current (normally 2:0 nA) is not reached when the charger is on (corona wire contaminated or grid separating corona and aerosol flow contaminated) | yes, critical error |  | 0 | 1 | 1 | idle
2 | Naneos Partector | 2 | 1 | 7 | 7 | 7 | High RH: the humidity sensor reports a value larger than 80% | only critical if humidity becomes condensing, 95% could be a better threshold |  | 1 | 0 | 2 | not used
2 | Naneos Partector | 3 | 1 | 2 | 2 | 9 | Electrometer offset high (> 5 or 10mV depending on instrument version) may indicate contamination of electrometer insulators |  | Data invalid, service reqd | 1 | 1 | 3 | calibration
2 | Naneos Partector | 4 | 1 | 1 | 2 | 9 | Flow low. The flow is calculated from auxiliary signals not measured directly. This may not be accurate as error message. | if flag is up, it must be critical | AM or NPL to verify flow with external flow meter
2 | Naneos Partector | 5 | 1 | 0 | 0 | 0 | Buffer overflow: internal data processing too slow to handle data may happen if writing to the SD-card is excessively slow (e.g. with a nearly full large SD card) | instrument might not be storing data, if critical wouldn't be any data, so no need to flag
2 | Naneos Partector | 6 | 1 | 0 | 0 | 0 | Generic error (no specifc condition) currently used for SD card missing | not using the SD card, okay to ignore
2 | Naneos Partector | 7 | 1 | 2 | 2 | 9 | Instrument calibrating (only PCB rev 2.3 and greater) | shouldn't occur, we don't calibrate the naneos (have no way to do so), if it did occur, would be critical
3 | LiCor LI-7200RS | 1 | 0 | 1 | 2 | 9 | Signal strength <= 86.71 % | yes critical
3 | LiCor LI-7200RS | 2 | 0 | 1 | 2 | 9 | Signal strength <= 73.37 % | yes critical
3 | LiCor LI-7200RS | 3 | 0 | 1 | 2 | 9 | Signal strength <= 46.69 % | yes critical
3 | LiCor LI-7200RS | 5 | 0 | 2 | 2 | 9 | PLL. Lock bit indicates that optical wheel is NOT rotating at the correct rate | yes critical
3 | LiCor LI-7200RS | 6 | 0 | 1 | 2 | 9 | Detector temperature not OK | yes critical because of infrared source
3 | LiCor LI-7200RS | 7 | 0 | 1 | 2 | 9 | Chopper wheel temperature not OK | less critical but still important
3 | LiCor LI-7200RS | 8 | 0 | 1 | 2 | 9 | Differential pressure not OK | suggest blockage or leak, critical in either case
3 | LiCor LI-7200RS | 9 | 0 | 2 | 0 | 0 | Aux input not OK | no aux inputs, ignore
3 | LiCor LI-7200RS | 10 | 0 | 1 | 1 | 2 | T inlet thermocouple not OK | in/out temps used to estimate cell temperature, data not critically affected (within other sources of instrument uncertainty) | this refers to the thermocouple itself not the temperature its reporting....not critical
3 | LiCor LI-7200RS | 11 | 0 | 1 | 1 | 2 | T outlet thermocouple not OK | in/out temps used to estimate cell temperature, data not critically affected (within other sources of instrument uncertainty) | this refers to the thermocouple itself not the temperature its reporting....not critical
3 | LiCor LI-7200RS | 12 | 1 | 0 | 0 | 0 | Sensor head type. 1 = LI-7200RS and LI-7200 | for information only
4 | Ecotech Serinus 40 | 9 | 1 | 2 | 2 | 9 | Cell temperature failure | yes critical | may not be critical
4 | Ecotech Serinus 40 | 10 | 1 | 2 | 2 | 9 | Cooler failure | yes critical | may not be critical
4 | Ecotech Serinus 40 | 11 | 1 | 2 | 2 | 9 | Converter failure | yes critical
4 | Ecotech Serinus 40 | 14 | 1 | 1 | 2 | 9 | Flow fault | yes critical
4 | Ecotech Serinus 40 | 15 | 1 | 1 | 0 | 0 | Any system error (the red instrument panel light is on) | if any status is bad, light will come on (includes warm up), if critical, it would be covered by other flags
4 | Ecotech Serinus 40 | 16 | 1 | 4 | 4 | 4 | Currently in warmup process | initial power up, data should be redacted by the instrument
4 | Ecotech Serinus 40 | 18 | 1 | 1 | 2 | 9 | Performing a background | on occasion for a few seconds but will hold the previous result, not actively measuring, should be redacted because we are mobile
4 | Ecotech Serinus 40 | 19 | 1 | 2 | 0 | 0 | Currently in span mode | shouldn't be unless at NPL
4 | Ecotech Serinus 40 | 20 | 1 | 1 | 0 | 0 | Currently in zero mode | shouldn't be unless at NPL
4 | Ecotech Serinus 40 | 21 | 1 | 2 | 2 | 9 | Instrument out of service (or in diagnostic mode or PTF compensation or control loop disabled or comms debugging enabled) | yes critical, instrument about to be serviced, could happen while driving if someone forgot to turn service mode off - can ignore if flag happened on road
4 | Ecotech Serinus 40 | 22 | 1 | 2 | 2 | 9 | High voltage failure | critical | 1 | 0
4 | Ecotech Serinus 40 | 23 | 1 | 2 | 2 | 9 | System power failure (not acutally possible to report) | critical, comes on after power has been restored, flags first reading | 0 | 0 | 0 | Measurement
5 | Magee Scientific AE33 | 0 | 1 | 2 | 2 | 9 | 1. Tape advance or 3. stopped (see bit 1) | data invalid during these time | 0 | 1 | 1 | Tape advance
5 | Magee Scientific AE33 | 1 | 1 | 2 | 2 | 9 | 2. First measurement - obtaining ATN0 or 3. stopped (see bit 0) | data invalid during these time | 1 | 0 | 2 | First measurement - obtaining ATN0
5 | Magee Scientific AE33 | 2 | 1 | 1 | 2019-01-02 | 2019-02-09 | Flow status low/high by more than 0.5 LPM | indicates > ~10% error, need to look at flow, if outside 4.5 LPM to 5.5 LPM range, then critical | 1 | 1 | 3 | Stopped
5 | Magee Scientific AE33 | 3 | 1 | 1 |  |  | Check flow status history |  | This is the second bit of the above flag, treat as first bit
5 | Magee Scientific AE33 | 4 | 1 | 1 | 2 | 9 | 1. Calibrating LED or 3. LED error (all channels calibration error) (see bit 5) | shouldn't see in normal operation, critical if seen
5 | Magee Scientific AE33 | 5 | 1 | 1 | 2 | 9 | 2. Calibration error (at least one channel OK) or 3. LED error (all channels calibration error) (see bit 4) | shouldn't see in normal operation, critical if seen
5 | Magee Scientific AE33 | 6 | 1 | 1 | 2 | 9 | Chamber status | shouldn't see in normal operation, critical if seen | 8 | 7 |  | Filter status | Flag
5 | Magee Scientific AE33 | 7 | 1 | 1 | 0 | 0 | 1. less than 30 filter spots left or 3. no filter left (see bit 8) | not critical, just a warning to change tapes, if no spots left, no data | 0 | 0 | 0 | OK | 0
5 | Magee Scientific AE33 | 8 | 1 | 1 | 0 | 0 | 2. less than 5 filter spots left or 3. no filter left (see bit 7) | not critical, just a warning to change tapes, if no spots left, no data | 0 | 1 | 1 | less than 30 spots left | 1
5 | Magee Scientific AE33 | 10 | 1 | 2 | 2 | 9 | 1. Stability test or 3. Change tape procedure or 5. Leakage test (see bits 11 and 12) | shouldn't see in normal operation, critical if seen | 1 | 0 | 2 | less than 5 spots less | 1
5 | Magee Scientific AE33 | 11 | 1 | 2 | 2 | 9 | 2. Clean air test or 3. Change tape procedure (see bit 10) | shouldn't see in normal operation, critical if seen | 1 | 1 | 3 | No filter left | 1
5 | Magee Scientific AE33 | 12 | 1 | 2 | 2 | 9 | 4. Optical test or 5. Leakage test (see bit 10) | shouldn't see in normal operation, critical if seen
5 | Magee Scientific AE33 | 14 | 1 | 2 | 2 | 9 | Auto clean air test status | shouldn't see in normal operation, critical if seen | 12 | 11 | 10
5 | Magee Scientific AE33 | 15 | 1 | 2 | 0 | 0 | CF card failure | maintenance only, data unaffected because we are taking data out via serial | 0 | 0 | 0 | 0 | OK
5 | Magee Scientific AE33 | 16 | 1 | 0 | 0 | 0 | Database size warning | maintenance only, data unaffected because we are taking data out via serial | 0 | 0 | 1 | 1 | Stability test
6 | 2BTechnologies Model 211-G | 18 | 1 | 2 | 2 | 9 | High scrubber temperature | yes critical, using a heated graphite scrubber to remove o3 and look at interference, can't take measurements without this | 0 | 1 | 0 | 2 | Clean air test
6 | 2BTechnologies Model 211-G | 19 | 1 | 2 | 2 | 9 | Low scrubber temperature | yes critical, using a heated graphite scrubber to remove o3 and look at interference, can't take measurements without this | 0 | 1 | 1 | 3 | Change tape procedure
6 | 2BTechnologies Model 211-G | 28 | 1 | 1 | 2/0 | 2019-02-09 | Cell B flow high | 1300cc/min | 1 | 0 | 0 | 4 | Optical test
6 | 2BTechnologies Model 211-G | 29 | 1 | 1 | 2/0 | 2019-02-09 | Cell B flow low | 900 cc/min | 1 | 0 | 1 | 5 | Leakage test
6 | 2BTechnologies Model 211-G | 30 | 1 | 1 | 2/0 | 2019-02-09 | Cell A flow high | 1300cc/min
6 | 2BTechnologies Model 211-G | 31 | 1 | 1 | 2/0 | 2019-02-09 | Cell A flow low | 900 cc/min | 19 | 18
7 | Aerodyne CAPSNO2 | 4 | 0 | 2 | 2 | 9 | Pump off, no filter | yes critical | 0 | 0 | 0 | OK
7 | Aerodyne CAPSNO2 | 4 | 2 | 2 | 2 | 9 | Pump off, filter in | yes critical | 0 | 1 | 1 | high scrubber temp
7 | Aerodyne CAPSNO2 | 3 | 1 | 2 | 2 | 9 | Baseline on - flush period | yes critical - mixture of street and zero air | 1 | 0 | 2 | low scrubber temp
7 | Aerodyne CAPSNO2 | 3 | 2 | 2 | 0 | 0 | Baseline on - measurement period | not representative of on road measurement, but is a proper measurement of a baseline | 1 | 1 | 3
7 | Aerodyne CAPSNO2 | 2 | 1 | 2 | 0 | 0 | LED is off (Used only for Pmssa monitor) | not applicable | Confirmed
Note on E1: 1 = requires human confirmation.
Note on F1: status flags 1 are rewritten as 0 or 2 signifying how they should be treated by a computer. 0 = publish, 2 = do not publish.